## Slide 1: Lepton Capability Analysis
Michael Hayes

## Slide 2: Previous Logging Software
Previously, captures were taken with ‘lepton_capture’ code Ben wrote
for(I=0;I<100;I++) lepton_capture <outpath.csv> <outpath.pgm> sleep; 5 done;
Each time a capture was taken, the program was run again and the connection had to be re-established- an output csv and pgm were then produced
Had ~3s overhead for each capture
During time between captures there was no communication between the pi and the lepton, so the lepton would begin to cool

## Slide 3: Changes to Logging Software
Edited a different piece of software Ben produced (lepton_calib)
Instead of re-establishing communication with the lepton every time a capture was taken, a constant stream was established and displayed
Added code to calculate and log the FPA temperature and the signal using various methods

## Slide 4: Different regions of interest used
C9 – 3x3 central pixels
C25 – 5x5 central pixels
D13 – 13 central pixels in ‘diamond’ orientation
Dxx – xx central pixels in ‘diamond’ orientation
R8 – all pixels in circle of radius 8
As signal values were recorded on the fly, a fixed radius was used instead of finding the size of the aperture as this would be more computationally intensive

## Slide 5: Temperature Evolution

## Slide 6: Pertubations
Lepton on average crashed around every 3 minutes during long term stability test
During this period, FPA temperature drops by around 0.2 degrees- the lepton then requires time to restabilise temperature
Crashes likely due to lepton hardware as opposed to Ben’s interface, however we should look in to ways of reducing frequency
There is not much indication a crash has occurred just by looking at screen- image appears to freeze very briefly then resumes as the lepton reboots
This likely will be a limiting factor as to the effectiveness of DFIRST in the hands of a clinician
Visual indication (green light etc) as to when camera is ready to take image? E.g. when abs(min-max) of last N FPA temperature readings is less than some threshold (~0.02 degrees)
Inaccuracy in regions before temperature has stabilised?
Chart: Long term stability

## Slide 7: Crashes- cont.

## Slide 8: Stability at ambient
Water circulator was removed and water in the jacket emptied; lepton was then left to run to assess temperature stability without influence of circulator
Water jacket still conducts heat away so not a perfect representation of how lepton will be used in DFIRST

## Slide 9: Stability outside jacket

## Slide 10: Calibration using new capture method

## Slide 11: Uncertainties in non-stable region
Chart: 

## Slide 12: Size of source effect using simulated foot

## Slide 13: Possible alternative calibration approach
Dx/Dz = …

## Slide 14: Misc
Polynomial does not appear to describe target temp well in non-constant temperature regime
Perturbations appear to affect FPA thermometer longer than microbolometer array- regions where signal is stable (to within std dev) but FPA temp is not
Resistor is measuring its own temperature as opposed to FPA’s
array able to dissipate heat quicker than sensor/is located closer to jacket?
Could predict where FPA perturbation will end and use this as temperature in polynomial
Chart: FPA temp
Chart: Signal